## Sheet1
User Need ID | User Need Summary | Design Input ID | Design Input Summary | Verification Test ID | Verification Method
UN-01 | Objectively measure elbow ROM | DI-01 | Measure ROM from 0° to 135° | VT-01 | Angle Verification Test
UN-01 | Objectively measure elbow ROM | DI-02 | Accuracy within ±5° | VT-02 | Accuracy Verification Test
UN-02 | Feedback on rehab performance | DI-03 | Display current angle | VT-03 | Functional Test
UN-03 | Track exercise repetitions | DI-04 | Count completed reps | VT-04 | Rep Counting Test
UN-05 | Monitor rehab progress | DI-05 | Record maximum angle | VT-05 | Session Metrics Test
UN-05 | Monitor rehab progress | DI-06 | Record minimum angle | VT-05 | Session Metrics Test
UN-05 | Monitor rehab progress | DI-07 | Record total repetitions | VT-05 | Session Metrics Test
UN-04 | Independent home use | DI-08 | Portable sensor configuration | VT-06 | Design Inspection
UN-06 | Easy-to-understand information | DI-09 | Present understandable data | VT-07 | Usability Review
UN-07 | Support natural arm movement | DI-10 | Allow natural arm movement | VT-08 | User Evaluation
UN-08 | Objective performance metrics | DI-11 | Provide reviewable data | VT-09 | Data Review Test
UN-09 | Compare ROM over time | DI-12 | Support multi-session comparison | VT-10 | Data Trend Review
UN-11 | Consistent measurements | DI-13 | Repeatable angle measurements | VT-11 | Repeatability Test
UN-12 | Accurate exercise representation | DI-14 | Sensor-based elbow position determination | VT-12 | Functional Test
UN-13 | Live rehabilitation measurements | DI-15 | Update displayed measurements during exercise | VT-13 | Functional Test