ROM-IQ

Design Inputs Document

Document Information

Document Title: Design Inputs Document

Project: ROM-IQ - Intelligent Range of Motion Monitoring System for Elbow Rehabilitation

Author: AK Gupte

Version: 1.0

Date: June 2026

1. Purpose

The purpose of this document is to translate user needs into measurable engineering requirements that guide the design, development, and verification of the ROM-IQ rehabilitation monitoring system.

2. Design Inputs

| ID | Related User Need | Design Input | Verification Method |
| --- | --- | --- | --- |
| DI-01 | UN-01 | The system shall measure elbow range of motion from 0° to 135°. | Angle Verification Test |
| DI-02 | UN-01, UN-12 | The system shall measure elbow range of motion with an error of ±5° compared to reference angles. | Accuracy Verification Test |
| DI-03 | UN-02, UN-13 | The system shall display current elbow angle measurements during an exercise session. | Functional Test |
| DI-04 | UN-03 | The system shall count completed elbow flexion-extension repetitions. | Rep Counting Test |
| DI-05 | UN-05 | The system shall record the maximum angle achieved during a rehabilitation session. | Session Metrics Test |
| DI-06 | UN-05 | The system shall record the minimum angle achieved during a rehabilitation session. | Session Metrics Test |
| DI-07 | UN-05, UN-10 | The system shall record the total number of repetitions completed during a rehabilitation session. | Session Metrics Test |
| DI-08 | UN-04 | The system shall operate using a portable microcontroller and sensor configuration suitable for home use. | Design Inspection |
| DI-09 | UN-06 | The system shall present rehabilitation measurements in a format that is understandable to users. | Usability Review |
| DI-10 | UN-07 | The system shall support elbow rehabilitation exercises while allowing natural arm movement. | User Evaluation |
| DI-11 | UN-08 | The system shall provide objective rehabilitation performance metrics that can be reviewed after testing. | Data Review Test |
| DI-12 | UN-09 | The system shall provide rehabilitation data that supports comparison of range of motion measurements over multiple sessions. | Data Trend Review |
| DI-13 | UN-11 | The system shall provide repeatable angle measurements under identical test conditions. | Repeatability Test |
| DI-14 | UN-12 | The system shall utilize sensor-based motion measurements to determine elbow position. | Functional Test |
| DI-15 | UN-13 | The system shall update displayed rehabilitation measurements during exercise performance. | Functional Test |

3. Design Input Review

The design inputs defined within this document establish the measurable engineering requirements for ROM-IQ. Each design input traces directly to one or more user needs and will be evaluated through verification activities during later phases of development.

4. Approval

This document serves as the baseline set of design inputs for the ROM-IQ development project.